## Fit
400 nm pump: energy vs Harmonics HWP, fit and working points (data 2026-10-28, pump throttle 16 deg)
Model: E_PM1 = c + A * cos^4(2 (HWP - HWPmax)). SHG energy goes as the square of the fundamental along the crystal axis, hence cos^4 (a free exponent fits best at exactly 4).
HWPmax (deg) | 49.65 | max energy; minimum is 45 deg below, at 4.65
A (uJ) | 16.6
c, floor (uJ) | 0.413 | NOT determined by the data (none below 23 deg): a leakage model, A (cos^2 + 0.048)^2, fits equally well (rms 0.038) with a floor of 0.06 uJ. Measure PM1 at 8 deg
Transmission PM1 -> IP | 0.636 | measured 2026-10-28
Spot FWHM at IP (um) | 227 | knife edge on the sample card, yaw set by eye
Peak fluence per uJ at IP (mJ/cm^2) |  | 4 ln2 / pi / FWHM^2 (incident, not absorbed)
Fit rms (uJ) | 0.036 | 17 points, 23-52 deg; max residual 0.08 uJ
Working points
Working point | HWP (deg) | PM1 (uJ) | IP (uJ) | Peak fluence (mJ/cm^2) | Note
YAG overlap (min energy) | 8 |  |  |  | Patrick: YAG at 8-10 deg. Energy here is NOT known from this data: 0.06-0.42 uJ at PM1 depending on the model. Measure PM1 at 8 deg
Bi time zero | 22 |  |  |  | Patrick's 'around here for Bi'. Robust: 2.10-2.16 uJ at PM1 across models (1 deg below the data)
Bi coarse t0 scan (optional) | 25 |  |  |  | Stronger signal for the wide first scan, if the sample tolerates it; measured point
Run 6 Bi angle (reference) | 16.245 |  |  |  | Old ued_setup value: now under 1 mJ/cm^2, no longer ~2 uJ
Max | 49.65
Table
HWP (deg) | Measured PM1 (uJ) | Fit PM1 (uJ) | Residual (uJ) | Fit IP (uJ) | Fit peak fluence (mJ/cm^2)
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23 | 2.5
24 | 2.95
25 | 3.4
26 | 3.95
27
28 | 5.1
29
30 | 6.4
31
32 | 7.7
33
34 | 9.3
35
36 | 10.8
37
38 | 12.2
39
40 | 13.6
41
42 | 14.7
43
44 | 15.8
45
46 | 16.5
47
48 | 16.9
49
50 | 17
51
52 | 16.8

## Data
Harmonics HWP (deg) | Energy Meter PM1 (uJ) | Copied from Energy_calibration_400nm_Ji_20261028.xlsx (pump throttle 16 deg).
52 | 16.8
50 | 17
48 | 16.9
46 | 16.5
44 | 15.8
42 | 14.7
40 | 13.6
38 | 12.2
36 | 10.8
34 | 9.3
32 | 7.7
30 | 6.4
28 | 5.1
26 | 3.95
25 | 3.4
24 | 2.95
23 | 2.5